# HemoSim edit doc - I16

HOW TO USE: edit the wording directly (Track Changes if you like). For layout or a figure (size, position, wrong/missing image) write a short instruction in [brackets] right where it applies. Send it back and I apply the wording to the page and the layout notes to the styling, then regenerate. This is a content+notes doc, so it will not look exactly like the web page.

# Measuring cardiac output

Cardiac output has appeared in every equation in this pathway. It is the flow term in all four interfaces and the common denominator that ties them into one circulation, as Module I1 argued. Every one of those arguments assumed the number was known. This module takes that assumption apart.

Module N7 T4 gave the working version: cold in the proximal port, a thermistor at the tip, output inversely proportional to the area under the washout curve, three measurements averaged at end expiration, unreliable in shunt and in tricuspid regurgitation. All of that is true and none of it is enough. A clinician who knows only that thermodilution is "unreliable" in tricuspid regurgitation cannot say whether the displayed 3.2 L/min is too high or too low, and has no basis for choosing what to do instead. The difference between the Novice treatment and this one is the difference between a list of conditions and the direction and size of the error each one produces.

Three methods are in routine use and they fail in unrelated ways, which is the most useful thing about having three. Thermodilution measures flow directly and assumes all the indicator you injected passes the detector once. Fick infers flow from an oxygen mass balance and assumes you know how much oxygen the patient is consuming. Doppler echocardiography measures a stroke volume geometrically and assumes a circular outflow tract whose diameter you measured correctly. When two agree the number is probably real, and when they disagree the disagreement localizes the broken assumption.

## Indicator dilution, and why the area under the curve is the answer

All dilution methods rest on one conservation statement. Put a known quantity of indicator into a stream, let it mix, and measure its concentration downstream until it has all gone past. What went in equals what came out, which is flow multiplied by concentration integrated over time.

m = Q × ∫C(t) dt   →   Q = m ÷ ∫C(t) dtm, quantity of indicator injected; Q, flow; C(t), indicator concentration at the detector as a function of time; ∫C(t) dt, the area under the concentration-time curve.

Read that inverse relationship physically. The same quantity of indicator must appear in the curve however fast the blood moves. A fast stream sweeps it past quickly, so concentration peaks early and falls fast and the area is small. A slow stream drags it past over a longer interval at a lower peak, and the area is large. Peak height tells you nothing on its own. Only the area carries the answer.

Thermodilution substitutes cold for a chemical indicator, which is what makes it practical: no blood is withdrawn and no arterial puncture is needed. Fegler established the thermal method in animals in 1954 and Ganz and colleagues brought it to the bedside in 1971, on the flow-directed balloon-tipped catheter Swan, Ganz and colleagues had described in 1970. The quantity of indicator is here a quantity of cold: injected volume multiplied by temperature difference, corrected for the different densities and specific heats of saline and blood. That is the Stewart-Hamilton relationship in the form the monitor computes.

Q = (V × (Tb − Ti) × K1 × K2) ÷ ∫ΔTb(t) dtQ, cardiac output; V, volume of injectate; Tb, blood temperature; Ti, injectate temperature; K1, a correction for the density and specific heat of injectate and of blood; K2, the computation constant, correcting for catheter dead space, warming of the injectate in transit down the catheter, and injection rate; ∫ΔTb(t) dt, the area under the temperature-time curve at the thermistor.

Everything that can go wrong is visible there. The numerator is what the monitor believes was injected, the denominator is what the thermistor actually saw, and the measurement is trustworthy exactly to the extent that those describe the same bolus.

K2 is the term clinicians enter and rarely think about, and it is not a fudge factor. Injectate does not arrive in the right atrium at the temperature it left the syringe: it warms travelling down a lumen surrounded by blood at 37 degrees, by an amount that depends on the length, wall thickness and volume of that lumen and on how fast the bolus is pushed. The constant is therefore specific to the catheter model, the injectate volume and the temperature range, which is why manufacturers publish a table rather than a number. Two consequences follow. Selecting the wrong catheter model biases every subsequent measurement, and nothing about the displayed curve looks abnormal. And the constant assumes a technique, since a slow injection gives the bolus more time to warm. Technique sits inside the equation, not next to it.

The integral runs from injection until the indicator has all passed, which never quite happens: some cooled blood returns through the systemic circuit and reaches the pulmonary artery a second time. Recirculation adds a slow hump to the tail, and integrating it would inflate the denominator and understate output. Every commercial computer therefore truncates the integration and extrapolates.

Integration begins at injection and terminates once the decay reaches roughly 30 percent of peak, after which the exponential is extrapolated to baseline. This is how the algorithm keeps recirculating indicator out of the measured area.Source: HemoSim hemodynamics module (v2022), slide 371.

[Layout note: figure shown small here. On the web it is currently capped at 640px wide. Say if you want it smaller/larger, moved, cropped, or replaced.]

That extrapolation is the hinge on which several error mechanisms below turn, because it assumes a single exponential generated by one pass of one bolus. Recirculation across a valve, a wandering baseline temperature or a second slug of cold arriving late all make the tail non-exponential, and the extrapolation has no way to know it is wrong. Hence the habit of reading the curve and not only the number beneath it.

Three curves, three messages. The normal curve peaks early and returns quickly, enclosing a small area. The low-output curve peaks later and lower and decays slowly, enclosing a large area that the equation converts into a low flow. The regurgitant curve is neither: it is a bolus redistributed in time by repeated crossings of the tricuspid valve, and its shape violates the single-pass assumption.Source: HemoSim hemodynamics module (v2022), slide 373.

[Layout note: figure shown small here. On the web it is currently full column width, about 750px. Say if you want it smaller/larger, moved, cropped, or replaced.]

## Technique is the dominant error term, and it is directional

Before any pathology is invoked, the largest source of scatter is the operator. That is a statement about the equation rather than a criticism: the numerator is a set of numbers typed into a monitor, the denominator is a physical event, and every discrepancy between them lands on the answer. The skill is not memorizing which errors read high but deriving the direction in the moment.

Work one. The monitor has been told 10 mL but only 8 mL is delivered. It computes the numerator with V equal to 10 mL, but only 80 percent of that cold arrived, so the thermistor records a shallower curve and the denominator is smaller. A fixed numerator over a smaller denominator gives a larger quotient: an under-delivered bolus reads falsely high. Every other technique error resolves the same way, by asking whether the monitor over- or underestimated the cold that reached the thermistor.

| Technique error | What it does to the equation | Measured cardiac output |
| --- | --- | --- |
| Less injectate delivered than programmed | Numerator held at programmed V while the measured area falls | Falsely high |
| More injectate delivered than programmed | More cold arrives than the numerator accounts for | Falsely low |
| Injectate warmer than programmed | Numerator overstates the thermal deficit and the real area is smaller | Falsely high |
| Injectate colder than programmed | More cold arrives than the numerator accounts for | Falsely low |
| Slow or hesitant injection | Cold is dispersed and partly absorbed before the thermistor, blunting the curve | Falsely high, and imprecise |
| Wrong catheter model, so wrong computation constant | K2 misstates warming in transit and dead space | Systematically wrong, direction set by the constant |
| Balloon left inflated | Flow past the thermistor is not the flow that received the bolus | Unpredictable, and unsafe |
| Rapid intravenous infusion running concurrently | Baseline pulmonary artery temperature drifts, so there is no stable baseline to integrate from | Unpredictable, sometimes grossly |

Three rows deserve expansion because the received wisdom about them is incomplete.

Iced versus room temperature injectate. The instinct is that colder is better, and a larger temperature difference does give a bigger signal, which matters most when the curve is shallow. But Stetz and colleagues, analysing fourteen published comparison studies, found no difference in measurement quality between rapid injection of iced and of room temperature indicator, and Nishikawa and Dohi recommend 10 mL of room temperature injectate for adults on the practical ground that iced injectate warms from the moment it leaves the ice bath. A controlled room temperature bolus beats an uncontrolled cold one.

Rapid concurrent volume infusion. Wetzel and Latson described major errors during rapid volume infusion, and the mechanism is baseline drift rather than anything to do with the bolus. The denominator is an area measured against a pulmonary artery temperature the algorithm assumes is stable. If cool fluid is pouring in through another line that baseline is moving, and an area computed against a sloping reference bears no defined relationship to flow. Pause the infusion, let the temperature settle, then measure.

What counts as a change. The convention of three to five injections with outliers discarded rests on a quantitative result. Stetz and colleagues concluded that with three measurements per determination, a difference of 12 to 15 percent, averaging 13 percent, is required before a change is clinically significant, rising to 20 to 26 percent for single measurements. An output moving from 4.0 to 4.4 L/min after an intervention has moved 10 percent, which is inside the noise of the method. Treating that as a response is treating measurement error.

[Physiologic Insight box: The precision of the reference method sets a ceiling on every comparison made against it. Critchley and Critchley, meta-analysing studies that compared cardiac output techniques with bias and precision statistics, showed that if a new device and thermodilution each carry a percentage error near 20 percent, the limits of agreement between them cannot be better than about 30 percent even if the new device is perfect, because the two errors combine. They proposed plus or minus 30 percent as the criterion for accepting a new method. So "validated against thermodilution" means a device agrees with an imperfect reference to within a tolerance widened to accommodate the reference's own imprecision. It is not a claim that either number is correct.]

## When the physiology, not the technique, breaks the measurement

Shunt produces two different errors that happen to point the same way, and conflating them causes confusion.

A left to right shunt does not break the measurement at all. The thermistor sits in the pulmonary artery, so it measures pulmonary blood flow, Qp, accurately. But pulmonary flow now exceeds systemic flow, Qs, by the shunt volume, and it is systemic output that determines oxygen delivery. Thermodilution has answered a real question, and it is the wrong question. It overestimates systemic output by the shunt flow.

A right to left shunt breaks the measurement itself. Some injectate crosses to the left heart and never passes the thermistor, so less indicator arrives than the numerator accounts for, the area shrinks, and the curve returns to baseline early, which the algorithm reads as fast washout. It also reads high, but by the mechanism of an under-delivered bolus rather than by measuring a different circuit. The remedies therefore differ. Right to left makes thermodilution unusable and the answer is Fick. Left to right makes it accurate but non-systemic, and the answer is to separate pulmonary from systemic flow using oximetry samples along the right heart to find the saturation step-up.

A regurgitant valve lets part of each bolus wash back into the right atrium and cross again later, so indicator arrives at the thermistor in installments rather than as a single pass. The curve becomes low, broad and long tailed, which is the grey curve above, and that shape is precisely what a single-exponential extrapolation was not built for.

First principles give two competing predictions, which is why the empirical literature is split. Spreading the same cold over a longer interval enlarges the true area, driving the output down. But if the smeared tail crosses the 30 percent termination threshold early, the extrapolation truncates a real part of the curve and drives it up.

The three studies worth knowing report exactly that spread. Cigarroa and colleagues (1989) paired thermodilution with Fick or indocyanine green in thirty patients. Without regurgitation, agreement was excellent, at 4.95 plus or minus 1.19 against 4.90 plus or minus 1.11 L/min. In the seventeen with regurgitation, thermodilution was consistently lower, at 4.22 plus or minus 1.45 against 4.99 plus or minus 1.67 L/min. That study is the origin of the underestimation teaching. Balik and colleagues (2002) compared thermodilution with simultaneous transoesophageal Doppler in twenty-seven ventilated patients by regurgitation grade, and the difference widened from 0.5 plus or minus 1.1 L/min with no or first degree, through 0.8 plus or minus 2.0 with second degree, to 1.9 plus or minus 2.3 with third degree, while correlation fell from 0.96 to 0.69. Note what widens fastest: the limits, not just the bias. Hoeper and colleagues (1999) made 105 paired thermodilution and Fick measurements in thirty-five patients with pulmonary hypertension and found a mean difference of plus 0.01 plus or minus 1.1 L/min, unaffected by severe regurgitation or by low output.

So the honest rule is not "thermodilution underestimates in tricuspid regurgitation". Severe regurgitation reliably degrades precision, while a systematic underestimate depends on the population. A single value in severe regurgitation therefore carries much wider uncertainty than usual, and direct Fick with a measured oxygen consumption is the arbiter when a second method disagrees.

Beat to beat variation in stroke volume is a precision problem, not a bias. A bolus injected during a run of short cycles samples a period of low stroke volume, one injected during long cycles samples the opposite, and neither is wrong about the interval it covered. The error is that the interval is short, so the response is to average more injections across more of the rhythm and to widen the least significant change accordingly. What atrial fibrillation does not do is produce a predictable directional error, so "the number is high because she is in atrial fibrillation" is not a legitimate inference in either direction.

Positive pressure ventilation makes right ventricular output vary within every breath, for the reasons Module I12 works through beat by beat, and a single measurement samples whatever part of that cycle it fell in.

Jansen and colleagues quantified this in ventilated pigs, taking fifty measurements at random points in the cycle at each of four levels of positive end-expiratory pressure. The maximum difference between randomly timed measurements was 40 percent on the left side of the heart and 70 percent on the right. Sorted by respiratory phase the variation proved systematic with only a small random component, which is the important part: the scatter is not noise, it is a reproducible waveform sampled at random. Their second finding is the uncomfortable one. On the left side end expiration was a satisfactory moment for estimating mean output, but on the right side, where a pulmonary artery catheter measures, no single moment held up across changing conditions, and applied positive end-expiratory pressure was one of the conditions that spoiled it.

Stevens and colleagues asked the clinical version in thirty-two patients, randomising injections to peak inspiration, end exhalation or random timing. Peak inspiration and end exhalation both gave much smaller standard deviations than random timing, and they recommended end exhalation. Nishikawa and Dohi recommend instead injecting at evenly spaced intervals across the ventilation cycle. Both are right and they answer different questions. End-expiratory injection is the most reproducible, always sampling the same point on the same waveform, which is what trending over hours needs. Evenly spaced injections come closest to the true cycle-averaged output, which is what an absolute value fed into a pulmonary vascular resistance or oxygen delivery calculation needs. What neither permits is injecting whenever the syringe is ready.

Everything above worsens as flow falls. A low output gives a shallow prolonged curve, so the signal is small relative to thermistor noise and baseline drift, more of the area falls in the extrapolated region, and the injectate spends longer in the right heart warming up. Low output is also where the number changes management most, which is a reason to corroborate rather than repeat.

[Offshoot / At-the-Bedside link: Look at the curves, not only the three numbers. A normal set is three curves of nearly identical shape: fast upstroke, single smooth peak, clean exponential decay to a flat baseline. A notched upstroke means the injection was not smooth. A tail that will not return to baseline, or a sloping baseline, means temperature drift, usually from a concurrent infusion. A low, broad, long-tailed curve in a patient whose perfusion looks adequate should prompt a look at the tricuspid valve before it prompts an inotrope. And before any of that, check the setup screen: the injectate volume the monitor believes it received and the catheter model it is using are the two most commonly wrong inputs in the whole measurement, and neither leaves a trace on the waveform.]

## The Fick principle

Fick's statement is a general mass balance: the uptake of a substance by the tissues equals blood flow multiplied by the arteriovenous difference in that substance's concentration. Applied to oxygen across the lung, the oxygen entering the blood each minute equals the flow through the lung multiplied by the oxygen each unit of blood picks up.

The mass balance made concrete. Blood arrives at the lung carrying roughly 150 mL of oxygen per litre and leaves with roughly 200 mL per litre, so each litre picks up about 50 mL. If the body consumes 250 mL per minute, five litres must pass through the lung each minute.Source: HemoSim hemodynamics module (v2022), slide 378, after Deranged Physiology.

[Layout note: figure shown small here. On the web it is currently capped at 640px wide. Say if you want it smaller/larger, moved, cropped, or replaced.]

CO = V̇O2 ÷ (CaO2 − CvO2)
CaO2 = (1.34 × Hb × SaO2) + (0.003 × PaO2)CO, cardiac output (L/min); V̇O2, oxygen consumption (mL/min); CaO2 and CvO2, arterial and mixed venous oxygen content (mL O2 per dL, computed identically from the respective saturation); Hb, hemoglobin (g/dL); SaO2, arterial saturation as a fraction; 1.34, mL of oxygen carried per g of saturated hemoglobin; 0.003, mL dissolved per dL per mmHg. Content is expressed per dL, so the difference must be multiplied by 10 to give mL per litre before dividing into V̇O2.

That last clause is where most arithmetic errors are born, so work an example through. Oxygen consumption 250 mL/min, hemoglobin 10 g/dL, arterial saturation 98 percent, mixed venous saturation 68 percent, ignoring dissolved oxygen, which contributes about 0.3 mL/dL and nearly cancels in the difference.

CaO2 = 1.34 × 10 × 0.98 = 13.1 mL O2/dL
CvO2 = 1.34 × 10 × 0.68 = 9.1 mL O2/dL
CaO2 − CvO2 = 4.0 mL O2/dL = 40.2 mL O2/L
CO = 250 ÷ 40.2 = 6.2 L/minAn arteriovenous content difference near 4 mL O2 per dL, or 40 mL per litre, is the normal value and is worth carrying as an anchor. It corresponds to an oxygen extraction ratio of roughly 30 percent.

Direct Fick measures oxygen consumption, by Douglas bag or metabolic cart, and with simultaneous arterial and mixed venous sampling it is the reference standard and the arbiter when thermodilution is in doubt. It is also slow, needs equipment most units do not keep at the bedside, and is corrupted by any leak in the circuit, which is why it is rarely done. Indirect Fick substitutes an assumed consumption from a population formula: 125 mL/min/m2 multiplied by body surface area, the LaFarge, Dehmer or Bergstra equations, or a flat 250 mL/min. This is what a catheterization report means by "Fick cardiac output" without qualification.

Narang and colleagues measured resting consumption by Douglas bag in 535 adults undergoing right heart catheterization and compared it with three published formulae. Measured consumption averaged 241 plus or minus 57 mL/min. All three formulae differed significantly from it, with median absolute differences of 28 to 38 mL/min, and estimated and measured values differed by more than 25 percent in 17 to 25 percent of patients. The discrepancy was worse in severe obesity, above a body mass index of 40, and was unaffected by sex or age.

Consumption is the numerator, so a proportional error in the estimate becomes the same proportional error in the output with no attenuation, and the direction is predictable from the patient rather than from the formula. A deeply sedated, paralysed, cooled, ventilated patient consumes less than the population estimate, because the work of breathing is abolished and skeletal muscle is silent, so the numerator is too big and indirect Fick overestimates. A febrile, agitated, shivering or thyrotoxic patient consumes more, so the numerator is too small and it underestimates. Severe pneumonia adds an error that degrades direct Fick too: inflamed lung consumes oxygen drawn from alveolar gas that never appears as a systemic arteriovenous difference, so uptake measured at the mouth exceeds what the content difference accounts for and the output reads too high. Several of these can coexist in one patient pulling opposite ways. That is the real case against indirect Fick in critical care: not that the formulae are badly fitted, but that the intensive care population is furthest from the resting adults they were fitted on.

The content difference sits in the denominator, so the fractional error in the output is roughly the fractional error in consumption plus the fractional error in the content difference. The second term behaves badly whenever the difference is small, because the same absolute sampling error is a larger fraction of it.

Work it. In the example above the content difference was 40.2 mL/L and the output 6.2 L/min. Misread the venous saturation by two points, at 70 rather than 68 percent, and the difference becomes 37.5 mL/L and the output 6.7 L/min, an error of 7 percent. Now take a hyperdynamic septic patient with the same hemoglobin and arterial saturation but a mixed venous saturation of 85 percent. The content difference is 17.4 mL/L and the output 14.4 L/min, and the same two-point error now gives 14.7 mL/L and 17.0 L/min, an error of 18 percent. Nothing changed about the sample. That is why Fick is least trustworthy in high-output states, and why an intracardiac shunt, an arteriovenous dialysis fistula or any other source of arterialized blood in the venous sample is so damaging: each raises the mixed venous saturation, narrows the denominator and inflates the output without touching the flow.

The mixed venous saturation must be a genuine mixture of everything returning to the heart, which is why it comes from the distal port in the pulmonary artery, downstream of the coronary sinus and both cavae, rather than from a central line. Module I11 takes up central venous oxygen saturation as a perfusion marker in its own right, and the two are not interchangeable. Two sampling faults raise the measured saturation and therefore inflate the output: aspirating too quickly draws oxygenated blood back across the pulmonary capillary bed, and sampling with the balloon inflated or from a catheter that has migrated into a permanent wedge draws pulmonary capillary blood outright. Draw slowly, balloon down, after confirming a pulmonary artery waveform.

## What happens when the two methods are compared

The disagreement between Fick and thermodilution looks contradictory across studies until you notice that they are not comparing the same thing. What decides the answer is whether oxygen consumption was measured or assumed.

Hoeper and colleagues, in the pulmonary hypertension study above, found a mean difference between thermodilution and Fick of plus 0.01 plus or minus 1.1 L/min across 105 paired measurements, holding even in severe tricuspid regurgitation and low output.

Opotowsky and colleagues asked the same question at scale against estimated Fick. In more than 15,000 adults undergoing right heart catheterization, 12,232 in the Veterans Affairs Clinical Assessment, Reporting and Tracking cohort and 3,391 at Vanderbilt, thermodilution and estimated Fick cardiac index correlated only modestly, at r equal to 0.65. Mean difference was negligible at minus 0.02 L/min/m2, but the 95 percent limits of agreement ran from minus 1.3 to plus 1.3 L/min/m2, which is minus 50 to plus 49 percent. The methods differed by more than 20 percent in 38 percent of patients.

The study then did something agreement statistics cannot: it asked which number was right, using survival as the referee. A low thermodilution index predicted mortality more strongly than a low estimated Fick index at both 90 days and one year. Where the two disagreed, a low thermodilution index with a normal estimated Fick index carried higher one-year mortality than normal by both methods, 15.4 against 12.3 percent, whereas a normal thermodilution index with a low estimated Fick index carried no excess risk at all.

So the correct statement is narrower than the one usually taught. Thermodilution and direct Fick agree closely. Thermodilution and estimated Fick do not, and where they diverge the fault lies with the assumed consumption rather than with the dilution curve. A report quoting a "Fick cardiac output" is not interpretable until you know which kind it was, and in most laboratories it is the estimated kind.

## Cardiac output by echocardiography

Echocardiography abandons both indicator dilution and oxygen balance and measures a volume geometrically. Picture the blood ejected through the left ventricular outflow tract during one systole as a cylinder. Its cross-sectional area is that of the outflow tract, and its length is the distance a parcel of blood travels during that systole, which is what a velocity integrated over the ejection period gives. Multiply for stroke volume, multiply by heart rate for output.

CSA = πr2 = π × (d ÷ 2)2
SV = CSA × VTI
CO = SV × HRCSA, cross-sectional area of the left ventricular outflow tract (cm2); d, outflow tract diameter (cm); r, its radius; VTI, velocity time integral, the area under the pulsed wave Doppler velocity envelope for one systole (cm); SV, stroke volume (mL, since 1 cm3 equals 1 mL); HR, heart rate.

The diameter is measured in the parasternal long axis view at mid systole with the aortic leaflets fully open, from the junction of the aortic leaflet with the septal endocardium to the junction with the anterior mitral leaflet posteriorly, inner edge to inner edge. It is one linear measurement, and it is then squared.

Squaring is the whole problem. An error of 0.2 cm sounds trivial on a structure 2 cm across, and it is a 10 percent error in the diameter. A diameter of 2.0 cm gives an area of 3.14 cm2, while 1.8 cm gives 2.54 cm2. With the same velocity time integral of 19 cm the stroke volumes are 59.7 and 48.3 mL, so at a heart rate of 100 the outputs are 6.0 and 4.8 L/min. Two millimetres of caliper placement has moved the same patient from a normal output to one that would prompt an inotrope, and no other single measurement in hemodynamics carries that leverage. So spend the time on the diameter, zoomed, on a frame chosen by scrolling rather than by luck, and having measured it once do not measure it again: the outflow tract does not change over the hours of a resuscitation, so re-measuring injects fresh random error into a series meant to show a trend.

Velocity is obtained with pulsed wave Doppler, not continuous wave, because pulsed wave interrogates a defined sample volume at a known depth rather than everything along the beam. The sample volume goes about 5 mm proximal to the aortic annulus in an apical five or three chamber view, and the target is a thin, hollow envelope. A thin band means every red cell in the sample volume is moving at nearly the same velocity, which is laminar flow and the condition under which one traced envelope represents the whole column. A filled-in, spectrally broadened envelope means the sample volume sits where velocities are dispersed, usually too close to the valve. Average three to five beats in sinus rhythm and five to ten in atrial fibrillation.

A traced outflow tract envelope, following the outer edge of the dense modal signal from the onset of ejection to aortic valve closure. The machine returns its integral in centimetres, and the thin hollow shape is the sign that the sample volume sits in laminar flow.Source: HemoSim hemodynamics module (v2022), slide 409.

[Layout note: figure shown small here. On the web it is currently capped at 640px wide. Say if you want it smaller/larger, moved, cropped, or replaced.]

Doppler measures only the component of velocity lying along the beam, so a beam not parallel to flow reports a velocity reduced by the cosine of the intercept angle. The error is therefore always in one direction: an off-axis beam underestimates velocity, the integral and the output, and never overestimates them. At 20 degrees the cosine is 0.94, so the output is understated by about 6 percent, which is the basis of the conventional 20 degree limit, and at 30 degrees the cosine is 0.87 and the understatement is 13 percent. Knowing the direction is what makes this actionable: an unexpectedly low echocardiographic output in a patient with a difficult window is a floor, not a measurement.

Alignment between the Doppler beam and the direction of flow. Because the reported velocity is the true velocity multiplied by the cosine of the intercept angle, misalignment always reduces the measured value.Source: HemoSim hemodynamics module (v2022), slide 408.

[Layout note: figure shown small here. On the web it is currently capped at 400px wide. Say if you want it smaller/larger, moved, cropped, or replaced.]

The two measurements that produce a cardiac output. The parasternal long axis image supplies the diameter, measured once. The apical five chamber pulsed wave trace supplies the velocity time integral, measured on every study.Source: HemoSim hemodynamics module (v2022), slide 411.

[Layout note: figure shown small here. On the web it is currently capped at 640px wide. Say if you want it smaller/larger, moved, cropped, or replaced.]

CSA = π × (2.0 ÷ 2)2 = 3.14 cm2
SV = 3.14 × 19 = 59.7 mL
CO = 59.7 × 100 = 5,970 mL/min = 5.97 L/minOutflow tract diameter 2.0 cm, velocity time integral 19 cm, heart rate 100. Because 1 cm3 equals 1 mL, the product of an area in cm2 and a length in cm is a volume in mL with no conversion factor.

The most useful move in bedside echocardiographic hemodynamics is to stop calculating cardiac output altogether. Cross-sectional area is a constant for a given patient, so it contributes nothing to any change in stroke volume. Every change in stroke volume is a change in the velocity time integral, and every change in output is a change in the integral or the heart rate. Blanco makes the case for tracking stroke volume with the integral itself, and output with the minute distance, defined as the integral multiplied by heart rate.

The gain is not just convenience. Dropping the area drops the only squared term, the only one whose error is amplified, so a tracked change is far more reliable than either absolute value it came from, and it works when the absolute value cannot be obtained at all. Blanco's cardiogenic shock example is exactly that case: a patient whose parasternal long axis window was too poor to measure a diameter, in whom the integral rose from 12.1 to 17.8 cm on dobutamine while heart rate rose from 104 to 111, giving a minute distance that rose from 1,248 to 1,976 cm/min, an increase of 58 percent. No cardiac output was ever calculated and the clinical question was fully answered. This is the logic Module I11 applies to the passive leg raise and the fluid challenge, where the response rather than the baseline is the measurement.

Mercado and colleagues compared 64 paired measurements in 38 ventilated intensive care patients with pulmonary artery catheters in place. Median bias was 0.2 L/min, limits of agreement minus 1.3 to plus 1.8 L/min, and percentage error 25 percent, inside the Critchley threshold, with a precision of 8 percent for the catheter and 9 percent for echocardiography. Tracking did better still: the four-quadrant concordance rate for changes in output was 94 percent. The other echocardiographic route, the biplane Simpson method, derives end-diastolic and end-systolic volumes by tracing the endocardial border in two orthogonal apical views and subtracting. It is more vulnerable to endocardial dropout, foreshortening and regional wall motion abnormality than the outflow tract method, and it is treated in the Expert pathway.

## Choosing a method, and validating a number before you treat it

Given three methods with unrelated failure modes, the choice is made by asking which assumption this patient breaks.

| Situation | Preferred method | Because |
| --- | --- | --- |
| Stable ventilated patient, catheter in place, trending wanted | Thermodilution, standardized to one point in the respiratory cycle | Needs no metabolic assumption, and a fixed injection phase minimises the least significant change |
| Right to left intracardiac shunt | Fick, ideally with measured consumption | Indicator leaves the circuit before the thermistor |
| Left to right intracardiac shunt | Oximetry step-up sampling to separate pulmonary from systemic flow | Thermodilution accurately measures pulmonary flow, which is not the flow that perfuses the body |
| Severe tricuspid regurgitation | Two methods, with direct Fick as arbiter | Precision is reliably degraded and the direction of bias is not predictable |
| Atrial fibrillation with a variable rate | More injections, or direct Fick | Fick averages over the whole sampling period and is indifferent to rhythm |
| Fever, shivering, agitation, changing sedation | Thermodilution or echocardiography, not indirect Fick | The assumed consumption is exactly the quantity that is unstable |
| Hyperdynamic sepsis, narrow arteriovenous difference | Thermodilution or echocardiography | A small denominator amplifies every sampling error in the Fick calculation |
| No catheter, and the question is whether an intervention worked | Outflow tract integral and minute distance | The squared term drops out of a tracked change and no invasive access is needed |

A cardiac output is a hypothesis until something independent agrees with it, and three cross-checks are available at the bedside.

The oxygen story. A low output should come with a low mixed venous saturation, because a body extracting the same oxygen from less flow returns blood with less left in it. An output of 2.2 L/min alongside a mixed venous saturation of 68 percent, a normal lactate, preserved urine output and warm extremities is not cardiogenic shock, it is a measurement to repeat, starting with the injectate volume on the setup screen. Conversely, an output of 3.0 L/min with a mixed venous saturation of 82 percent and a lactate of 5.0 mmol/L is not an artefact to explain away: it is the signature of impaired extraction in distributive shock. Module I6 gives the mechanism and Module I11 the markers.

The rest of the panel. Divide the output by heart rate and ask whether the implied stroke volume is plausible for the ventricle you just looked at. Put it through the systemic vascular resistance calculation and ask whether the resistance matches the vasopressor requirement and the pulse pressure. A resistance wildly at odds with the phenotype usually means the flow term is wrong, since it is the least directly measured quantity in that expression.

The magnitude of change. Before concluding an intervention worked, check that the change exceeds the least significant change for the method: roughly 13 percent for triplicate thermodilution, 22 percent for a single measurement, and 10 to 15 percent for a tracked integral. Anything smaller is a coin toss dressed as data.

[Offshoot / At-the-Bedside link: When a cardiac output disagrees with the patient, the order of operations is fixed. Check the inputs first: injectate volume, injectate temperature, catheter model on the monitor, balloon down, no rapid infusion running. Then look at the curves for shape, not just the averaged number. Then ask what the anatomy is doing, meaning the tricuspid valve, the septum and the rhythm. Only then reach for a second method, and pick the one whose assumptions this patient does not break. Treating a number that has failed all three checks is the most avoidable harm attributable to hemodynamic monitoring, because the catheter does not cause the harm, the inotrope started on the strength of a bad number does.]

## What a measured flow is worth

One family of methods has been left aside. Minimally invasive devices infer output from the arterial pressure waveform, transpulmonary indicator dilution or thoracic bioreactance, and continuous pulmonary artery catheters warm blood with a thermal filament rather than injecting a bolus. Each carries its own assumptions about arterial compliance, waveform morphology or thoracic impedance, and therefore its own list of patients in whom it should not be believed. Those are treated at equation depth in the Expert pathway.

The three methods taught here are not three attempts at the same measurement. Thermodilution counts a quantity of cold and assumes it all passed the thermistor once. Fick counts oxygen and assumes you know how much the patient is burning. Doppler measures a column of blood and assumes a circle whose diameter you got right. Each is accurate for the question it actually answers, and every failure described here is a case where that was not the question being asked.

That is why measurement belongs at the end of this sequence rather than the beginning. Module I3 established that a displayed pressure is an inference from an apparatus, and this module makes the same case for a displayed flow. Module I15 did it for the wedge, whose assumption chain fails link by link in much the same way. The common discipline is to hold every displayed number as a claim with conditions attached, and to check them before it changes what you do. Module I17 takes whole patients and asks you to do exactly that, with the interfaces, the waveforms and the flow measurements all in play at once.

## References

To DROP a reference, strike it. To ADD one, write [add: author, year, journal, roughly what it is] rather than typing the citation out. Do not correct the text of an entry here: every author list, volume, issue, page range, year and DOI is fetched from PubMed when the page is rebuilt, so an edit made here is discarded and a typo made here is caught. This list is what the page currently carries.

1. Fegler G. Measurement of cardiac output in anaesthetized animals by a thermodilution method. Q J Exp Physiol Cogn Med Sci. 1954;39(3):153-164. doi:10.1113/expphysiol.1954.sp001067.

2. Swan HJ, Ganz W, Forrester J, Marcus H, Diamond G, Chonette D. Catheterization of the heart in man with use of a flow-directed balloon-tipped catheter. N Engl J Med. 1970;283(9):447-451. doi:10.1056/NEJM197008272830902.

3. Ganz W, Donoso R, Marcus HS, Forrester JS, Swan HJ. A new technique for measurement of cardiac output by thermodilution in man. Am J Cardiol. 1971;27(4):392-396. doi:10.1016/0002-9149(71)90436-x.

4. Forrester JS, Ganz W, Diamond G, McHugh T, Chonette DW, Swan HJ. Thermodilution cardiac output determination with a single flow-directed catheter. Am Heart J. 1972;83(3):306-311. doi:10.1016/0002-8703(72)90429-2.

5. Nishikawa T, Dohi S. Errors in the measurement of cardiac output by thermodilution. Can J Anaesth. 1993;40(2):142-153. doi:10.1007/BF03011312.

6. Stetz CW, Miller RG, Kelly GE, Raffin TA. Reliability of the thermodilution method in the determination of cardiac output in clinical practice. Am Rev Respir Dis. 1982;126(6):1001-1004. doi:10.1164/arrd.1982.126.6.1001.

7. Stevens JH, Raffin TA, Mihm FG, Rosenthal MH, Stetz CW. Thermodilution cardiac output measurement. Effects of the respiratory cycle on its reproducibility. JAMA. 1985;253(15):2240-2242. doi:10.1001/jama.253.15.2240.

8. Jansen JR, Schreuder JJ, Bogaard JM, van Rooyen W, Versprille A. Thermodilution technique for measurement of cardiac output during artificial ventilation. J Appl Physiol Respir Environ Exerc Physiol. 1981;51(3):584-591. doi:10.1152/jappl.1981.51.3.584.

9. Wetzel RC, Latson TW. Major errors in thermodilution cardiac output measurement during rapid volume infusion. Anesthesiology. 1985;62(5):684-687. doi:10.1097/00000542-198505000-00035.

10. Cigarroa RG, Lange RA, Williams RH, Bedotto JB, Hillis LD. Underestimation of cardiac output by thermodilution in patients with tricuspid regurgitation. Am J Med. 1989;86(4):417-420. doi:10.1016/0002-9343(89)90339-2.

11. Balik M, Pachl J, Hendl J. Effect of the degree of tricuspid regurgitation on cardiac output measurements by thermodilution. Intensive Care Med. 2002;28(8):1117-1121. doi:10.1007/s00134-002-1352-0.

12. Hoeper MM, Maier R, Tongers J, Niedermeyer J, Hohlfeld JM, Hamm M, Fabel H. Determination of cardiac output by the Fick method, thermodilution, and acetylene rebreathing in pulmonary hypertension. Am J Respir Crit Care Med. 1999;160(2):535-541. doi:10.1164/ajrccm.160.2.9811062.

13. Narang N, Thibodeau JT, Levine BD, Gore MO, Ayers CR, Lange RA, Cigarroa JE, Turer AT, de Lemos JA, McGuire DK. Inaccuracy of estimated resting oxygen uptake in the clinical setting. Circulation. 2014;129(2):203-210. doi:10.1161/CIRCULATIONAHA.113.003334.

14. Opotowsky AR, Hess E, Maron BA, Brittain EL, Barón AE, Maddox TM, Alshawabkeh LI, Wertheim BM, Xu M, Assad TR, Rich JD, Choudhary G, Tedford RJ. Thermodilution vs estimated Fick cardiac output measurement in clinical practice: an analysis of mortality from the Veterans Affairs Clinical Assessment, Reporting, and Tracking (VA CART) program and Vanderbilt University. JAMA Cardiol. 2017;2(10):1090-1099. doi:10.1001/jamacardio.2017.2945. Free full text.

15. Critchley LA, Critchley JA. A meta-analysis of studies using bias and precision statistics to compare cardiac output measurement techniques. J Clin Monit Comput. 1999;15(2):85-91. doi:10.1023/a:1009982611386.

16. Mercado P, Maizel J, Beyls C, Titeca-Beauport D, Joris M, Kontar L, Riviere A, Bonef O, Soupison T, Tribouilloy C, de Cagny B, Slama M. Transthoracic echocardiography: an accurate and precise method for estimating cardiac output in the critically ill patient. Crit Care. 2017;21(1):136. doi:10.1186/s13054-017-1737-7. Free full text.

17. Blanco P. Rationale for using the velocity-time integral and the minute distance for assessing the stroke volume and cardiac output in point-of-care settings. Ultrasound J. 2020;12(1):21. doi:10.1186/s13089-020-00170-x. Free full text.